Case 01: Cranial Plate Design

Design and Development of a Patient-Specific Cranial Plate and Mold Using Medical Imaging and Additive Manufacturing Techniques

1. Introduction

1.1 Cranial Defect

Cranial defects are discontinuities in the skull caused by trauma, decompressive craniectomy, tumor resection, congenital abnormalities, or infections. These defects can compromise brain protection, aesthetics, and patient quality of life. Restoration of the cranial vault is achieved through a surgical procedure known as cranioplasty. PMMA is commonly used because of its low cost, biocompatibility, and ease of fabrication.

1.2 Need for Patient-Specific Implant (PSI)

Conventional manually shaped implants often result in poor anatomical fit and longer surgical time. Patient-Specific Implants (PSIs) are designed directly from CT images and accurately match the patient's anatomy, leading to improved cosmetic and functional outcomes.

2. Objective

The objective of this case study is:

- To design a patient-specific cranial implant from CT data.

- To reconstruct the skull defect using mirror anatomy.

- To generate a mold for PMMA implant fabrication.

- To demonstrate a low-cost additive manufacturing workflow.

- Manufacturing of Titanium and Peek Plate

3. Materials and Software Used

| Component | Software / Material |
| --- | --- |
| Medical Imaging | CT Scan (DICOM) |
| Segmentation | D2P / 3D Slicer |
| 3D Reconstruction | D2P |
| Implant Design | Autodesk Meshmixer |
| Slicing | Cura |
| Additive Manufacturing Material | Ti-6Al-4V, PEEK.PLA |
| Printing Technology | SLM (Titanium), FDM/SLS (PEEK) |

Approaches:

1. Patient-Specific PMMA Cranial Implant Using Mold-Based Manufacturing Approach:

Method:

STL Skull Model Import

↓

Mesh Repair (Inspector)

↓

Defect Boundary Selection

↓

Symmetry-Based Mirror Selection

↓

Boundary Smoothing

↓

Surface Separation

↓

Mirror Reconstruction

↓

Thickness Generation (3–5 mm)

↓

1 mm Skull Offset

↓

Boolean Difference

↓

Patient-Specific Cranial Plate Creation

↓

STL Export

Result

B) Design Specifications of Titanium PSI Cranial Implant

| Parameter | Specification |
| --- | --- |
| Implant Type | Patient-Specific Cranial Implant |
| Design Method | Mirror Reconstruction |
| Material | Ti-6Al-4V ELI |
| Implant Thickness | 3–5 mm |
| Central Region | Circular Perforations |
| Edge Region | Gyroid Lattice |
| Gyroid Width | 20–40 mm (2–4 cm) |
| Purpose of Gyroid | Bone Ingrowth & Osseointegration |
| Fixation Method | Titanium Screws |
| File Format | STL |

Design Specifications of PEEK PSI Cranial Implant

| Parameter | Specification |
| --- | --- |
| Implant Type | Patient-Specific Cranial Implant |
| Design Method | Mirror Reconstruction |
| Material | Medical Grade PEEK |
| Implant Thickness | 2.5–4.0 mm |
| Central Region | Circular Perforations |
| Edge Region | Gyroid Lattice |
| Gyroid Width | 10–20 mm |
| Purpose of Gyroid | Tissue Integration & Weight Reduction |
| Fixation Method | Titanium Cranial Screws |
| File Format | STL |